- Technical Sheet

Click here to email us, or call (972) 961-4655 to speak to a sales representative.

# Hospira/Abbott Aim+ Ambulatory Pump # Refurbished, 1 year warranty Pump Specifications. Dimensions:; 6"H x 4"W x 1.5"D. Weight:; 1 lbs. Power Sources:; Wall Plug-in AC Power Supply, 120 VAC, 60 Hz, 0.06A. Battery:; Four AA Alkaline Batteries. Battery Pack:; Rechargeable using AC Power Supply. Power Capacity:; Four AA Alkaline Batteries or Battery Pack, at least 3400 mL at rates above 50 mL/hr and at least four days at rates below 10 mL/hr. Pump Mechanism:; Wall Plug-in AC Power Supply, 120 VAC, 60 Hz, 0.06A. Memory Protection:; Current Program and 512 event history log protected by internal battery-backed memory for at least one year after power is removed. System Accuracy:; ±5% nominal. TPN Program Specifications. Volume:; 1 ml - 9600 ml. Total Infusion Time:; 1 min - 24 hours. Taper Up/Down Times:; 1 min - 3 hours. KVO Rate:; 1.0 ml/hr - 5.0 ml/hr. KVO Time:; 1 min - 24 hours. Container Size:; TPN plus KVO Volume - 9999 ml. Pain Management Specifications. Concentration:; 0.1 mg/ml - 50 mg/ml, or 1 µg/ml - 1000 µg/ml. Delivery Rate:; 0.0 ml/hr continuous plus bolus or 0.1 ml/hr continuous only - 5.0 ml/hr subq or 25.0 ml/hr epidural, or intravenous 0.0 mg/hr - 1250 mg/hr, or 1 µg/hr - 25000 µg/hr. Loading and Bolus Doses:; 0.0 ml/hr continuous plus bolus or 0.1 ml/hr continuous only - 5.0 ml/hr subq or 25.0 ml/hr epidural, or intravenous 0.0 mg - 1250 mg, or 1 µg - 25000 µg. Loading & Bolus Doses Rate:; approximately 125 ml/hr. Bolus Lockout:; 5 min - 999 min. Boluses Per Hour Limit:; 1 per hour - 12 per hour. Container Size:; 0.1 ml - 9999.9 ml 0.1 mg - 9999.9 mg, or 1 µg - 999999 µg. Intermittent Delivery Specifications. Concentration:; 0.1 mg/ml - 1000.0 mg/ml, or 1 µg/ml - 1000 µg/ml. Dose Amount:; 0.1 ml - 9600 ml 0.1 mg - 9900.0 mg, or 1 µg - 999900 µg. Infusion Time Per Dose:; 1 min - 24 hours. Dose Frequency:; 1 min - 96 hours. Number of Doses in Container:; 1 dose - 999 doses. KVO Rate:; 0.1 ml/hr - 5.0 ml/hr 0.1 mg/hr - 5000 mg/hr, or 1 µg/hr - 5000 µg/hr. Container Size:; 0.1 ml - 9999.9 ml 0.1 mg - 9900.0 mg, or 1 µg - 999900 µg. Variable Time Delivery Specifications. Concentration:; 0.1 mg/ml - 1000 mg/ml, or 1 µg/ml - 1000 µg/ml. Base Start/Stop Times:; 12:00AM - 11:59PM. Base Rate:; 0.1 ml/hr - 400 ml/hr 0.1 mg/hr - 9900 mg/hr, or 1 µg/hr - 400000 µg/hr. Daily Phase Amount:; 0.1 ml - 9600 ml 0.1 mg - 9900.0 mg, or 1 µg - 999900 µg. Phase Start/Stop Times:; 12:00AM - 11:59PM. Phase Percentage:; 0.1% - 100%. Phase Rate:; 0.1 ml/hr - 400 ml/hr 0.1 mg/hr - 9900 mg/hr, or 1 µg/hr - 400000 µg/hr. Phase Dose:; 0.1 ml - 9600 ml 0.1 mg - 9900.0 mg, or 1 µg - 999900 µg. Bolus Dose:; 0.1 ml - 25 ml 0.1 mg - 9900.0 mg, or 1 µg - 25000 µg. Bolus Dose Delivery Rate:; Approximately 400 ml/hr. Bolus Lockout:; 5 min - 999 min. KVO Rate:; 0.1 ml/hr - 5 ml/hr 0.1 mg/hr - 5000 mg/hr, or 1 µg/hr - 5000 µg/hr. Container Size:; 0.1 ml - 9999.9 ml 0.1 mg - 9900.0 mg, or 1 µg - 999900 µg. Continuous Delivery Specifications. Concentration:; 0.1 mg/ml - 1000 mg/ml, or 1 µg/ml - 1000 µg/ml. Rate:; 0.1 ml/hr - 400 ml/hr 0.1 mg/hr - 9900 mg/hr, or 1 µg/hr - 400000 µg/hr. Container Size:; 0.1 ml - 9999.9 ml 0.1 mg - 9900 mg, or 1µg - 999900 µg.