Baxter Colleague CXE Specs:

| Ac power Requirements | 100-120 VAC, 60Hz |
| --- | --- |
| Leakage Current | Less than 300 u amps earth leakage (tested per UL 2601) |
| External Fuses | 1.6 amp Type T (time delay) 250V |
| Power Cord | 2.9 m (9 feet) long with integrally molded hospital grade plug |
| Battery Supply System | Colleague 3 Pump - 3 hours operating time at 100 mL/hr for all three channels (fully charged new batteries) 12 hours to recharge to 80% of capacity Internal charge system recharges batteries whenever the device is connected to an AC outlet |
| Range of Programmable Flow Rates | Primary Infusion on each pump channel 0.1 to 99.9 mL/hr in 0.1 mL/hr increments 1 to 1200 mL/hr in 1 mL/hr increments Piggyback Infusion on each pump channel 0.1. to 99.9 mL/hr in 0.1 mL/hr increments 1 to 500 mL/hr in 1 mL/hr increments Note: Rate limits can be configured for values less Than those stated above. |
| Volume to be Infused | 0.1 - 99.9 mL in 0.1 mL increments 1- 9999 mL mL/hr in 1 mL increments Note: Volume to be Infused limits can be configured for values less than those stated above |
| KVO | 0.1 - 5 mL/hr in 0.1 mL increments (configurable option) or programmed rate whichever is less. |
| Priming rate | 500 mL/hr (to be used only when not connected to patient) |
| Advance Air Rate to Advanced an Air Bubbles | At the same rate programmed for the current primary or piggyback infusion. |
| Air Bubble Setting | Measured accumulations of approximately 25 microliters 50 microliters 100 microliters 150 microliters The Air Bubble Setting is a configurable option. |
| Auto Restart | Allows the pump channel to automatically restart if an occlusion is relieved within one minute after detection. This configuration option allows 0 to 9 restarts before manual intervention is required. |
| Nominal Downstream | Alarm sensitivity is a configurable option. Rate rangein mL/hr <21; 21- 200; >200. 2 psig; 4 psig; 6 psig; Minimum. 5 psig; 8 psig; 11 psig; Moderate. 9 psig; 12 psig; 15 psig; Maximum. |
| Device Size | Approximately 13.88" H x 7.75" W x 8.00" D (353 x 197 x 203 mm) Not including clamp knob or power cord |
| Device weight | Approximately 17.5 lbs (7.9 kg) |
| Environmental Operating Limits | 59º F - 100º F (15º C - 38º C) |
| Environmental Storage and Transport Limits | -20º F - 135º F (-29ºC - 57º C), 10-100% Reliable Humidity including condensation, 500-1060 hPa Barometric Pressure |
| Accessories | Configuration Transfer Cable (Product code 2M8155) |
| Ground Impedance | Less than 0.2 ohms (tested per UL-2601) |
| Solution Container Height | Pump will operate at all programmable flow rates with a solution container height ranging from -36 to +48 inches. Container height is measured from the pump channel selected for use. |
| Interfaces | Attachment to a remote nurse call system A variety of systems exist which enable alert and alarm conditions to be transmitted to a remote location via the communications port on the rear panel assembly of the pump. See the service manual for additional information. |
| ADMINISTRATION SETS | Baxter sets equipped with keyed On/Off clamps. 2C5331, 2C5417, 1C8160, 2C5431, 2C5418 |